## Slide 1: Single–axis Accelerometer Performance TIA Optimization
Presented by
Maxwell McFarlane & Mohammad Abdelnaby
Supervised by Prof. Jérôme Juillard
ICS905
High performance Sensors & Transducers (HiPStr)

## Slide 2: Agenda 01 02 04 03
TIA Stage Expectations
Top Level
TIA Stage LTSpice Simulations
TIA Performance
05
06
Post-TIA Process
Conclusion

## Slide 3: Top Level - Design Model
Sensor Properties | 
Differential Capacitance Sensitivity (α) | 2.35
Positional Sensitivity (β) | 4.1
1% Bandwidth | DC to 270 Hz
1% Dynamic Range | +/- 55
1% Temperature Range | +/- 40 °C
Differential Capacitance Noise Density | 1.2
Precision (270Hz BW) | 0.008
Sensor Properties | 
Differential Capacitance Sensitivity (α) | 
Positional Sensitivity (β) | 
1% Bandwidth | DC to 270 Hz
1% Dynamic Range | 
1% Temperature Range | +/- 40 °C
Differential Capacitance Noise Density | 
Precision (270Hz BW) | 
Amplifier Properties | Constraints
 | 100
 | 1pF
Bias Voltage Amplitude () | 5 V
Bias Frequency ) | 10 1 M
Parasitic Capacitance ) | 25 p
Amplifier Properties | Constraints
Sensor System
Electrical System
Acceleration noise
(thermal)
Current noise
(thermal, shot, flicker)
Quantizer noise

## Slide 4: Top Level - Maxwell Analogy
Representing Sensor System as electrical system:
  Equation of Motion →
  Maxwell Analogy () →
  Motional Components Equivalence →
Maxwell’s Analogy allows conversion to electrical element
Sensor Output Current is
Mechanical domain | Electrical domain
Velocity () | Current ()
Force () | Voltage ()
Mass () | Inductance ()
Damper () | Resistance ()
Compliance () | Capacitance ()
Mechanical domain | Electrical domain
Norton Model Used for LT Simulation
Schematic Representation of Norton Model
Capacitance-Displacement Dependence
  →
Bias Amplitude and Frequency
are proportional to input current

## Slide 5: Top Level - Sensor Model LTSpice Response
Figure 1 : Gain Response of Sensor Output (Value @ 0 – 270, Compare output 1% Bandwidth [Check if Correct], …

## Slide 6: TIA Stage – Expectations/Approach
Low Noise is priority!
  Potential source : OpAmp input voltage/current noise & feedback resistor
  OPamp Total Current Noise + +
  TIA will contribute the most noise to sensor!
Gain Response determines: Bandwidth & SNR
  GBP will determine OPamp bandwidth
  Gain Linearity! impacts sensor 1:1 translation
Found 2 OPamps that matched criteria:
  Linear Technology LTC6268 500MHz Ultra-Low Bias Current FET Input Op Amp
    GBP = 500 MHz, Input Voltage Noise = 4.9 nV/√Hz, Input Current Noise = 5.6 fA/√Hz
  Analog Devices ADA4528-1 Precision, Ultralow Noise, RRIO, Zero-Drift Op Amp
    GBP = 3 MHz, Input Voltage Noise = 5.6 nV/√Hz, Input Current Noise = 0.7 pA/√Hz
Quick Estimation! Since,
Assuming, Vbac = 5V and Vbf = 10kHz :
LTC6268 has much higher GBP and lower current noise
Expect better performance from this opamp
Notes: If we consider our TIA stage as an ”LNA” amplifier for our system. We know that this amplifier has the largest impact on our noise figure until the ADC at the receiver.

## Slide 7: TIA Stage – LTSpice TestBench Schematic
Voltage Source used to plot
TIA current
Param. Settings
Measurement
Instructions
Sweep/Step
Instructions
Acceleration source w/ Noise
Biasing Source
Supply Voltages
LTC6268 OPAMP
ADA4528-1 OPAMP
Sensor Circuit

## Slide 8: TIA Stage – LTSpice – (.ac) Gain (Cf Sweep)
As Cf increases, bandwidth is decreased. Best to choose
Cf determines linearity at our bias frequency
Gain 10dB delta over bias frequency range!

## Slide 9: TIA Stage – LTSpice – (.ac) Gain (Rf Sweep)
As Rf increases, Gain Increases
Take into account linearity around center frequency!!
-12.5 mdB delta (0.14%) @ 10KHz
-381 mdB delta (4.48%) @ 100KHz
Most linear region is around bottom of bias frequency range
270Hz
Ideal TIA Gain Properties | Value
Ideal TIA Gain Properties | Value
Notes: We had room to choose a higher feedback resistance but in order to reduce noise at later stages during noise analysis we lowered noise as much as possible here.
For our given values we, see that staying in the 10KHz bias frequency gives us a linear output gain and a

## Slide 10: TIA Stage – LTSpice – (.trans) Accel Sweep
Sanity Check!
Acceleration translates to higher current!
In slide 6, expected
Instead,
Difference maybe caused by sensor implementation

## Slide 11: TIA Stage – LTSpice – (.trans) Vbac Sweep
Sanity Check!
Higher Vbias Amplitude results in higher current!
Based on expected current & gain,
We expected ,

## Slide 12: TIA Stage – LTSpice – (.trans) Vbf Sweep
Sanity Check!
Higher Vbias frequency means greater input current!
The output of the ADA4528-1 has a response inverse to expectation!
Notes: Higher frequency kills the output voltage/current.
Could likely contribute this to the GBP being near 3M where the LTC6268 has a GBP of 50M
Or Seeing a change in comparison between gains between op amps in transient sims. Consider this is probably do to input capacitances that are a 10x difference

## Slide 13: TIA Stage – LTSpice – (.trans)Vbf Sweep
A clearer view that the higher bias frequency
is proportional to higher output current.

## Slide 14: TIA Stage – LTSpice – (.noise) VRf Sweep
Expected Precision of LTC6268 @ Vbf = 10k Rf = 1M Cf = 1pF | Value
Expected Precision of LTC6268 @ Vbf = 10k Rf = 1M Cf = 1pF | Value
Sanity Check!
Seeing 3x contribution to noise at output
@ Rf is 1Meg, noise is all lower at 10k than the 100KHz operating frequency!
Notes: Lower noise because of lower gain at the output
Though we expected lower current noise due to the resistor the gain contributes most to the output noise

## Slide 15: TIA Performance
TIA Performance | Value @ Rf = 100 Cf = 12p | Value @ Rf = 5 Cf = 12p | Value @ Rf = 1M Cf = 12p
SNR Range [55-1] | [148 dB, 99 dB] | [143 dB, 94 dB] | [138 dB, 89.3 dB]
Differential Capacitance Noise Density |  |  | 
(*Includes ) |  |  | 
TIA NF | 1.15dB | 1.94dB | 3.07dB
TIA Diff. Cap. Noise Density |  |  | 
TIA Diff. Cap. Noise | 4.45 aF | 7.74 aF | 13.2 aF
TIA Acceleration Precision Contribut. | 0.0009 | 0.0017 | 0.0028
TIA Performance |  |  | Value @ Rf = 1M Cf = 12p
 | [148 dB, 99 dB] | [143 dB, 94 dB] | [138 dB, 89.3 dB]
Differential Capacitance Noise Density |  |  | 
TIA NF | 1.15dB | 1.94dB | 3.07dB
TIA Diff. Cap. Noise Density |  |  | 
TIA Diff. Cap. Noise | 4.45 aF | 7.74 aF | 13.2 aF
TIA Acceleration Precision Contribut. |  |  | 
Acceleration noise
(thermal)
Current noise
(thermal, shot, flicker)
Quantizer noise
Even though the output noise is higher at the output the SNR is best at the highest gain possible
Precision of the system including
The sensor and the TIA together.
Convert Sensor noise to noise at the
Output of TIA

## Slide 16: Post-TIA Process
Is a 2nd Gain Stage necessary? No!
  Expected TIA Voltage Range : [5.58V
  Typical ADCs FS (Typ. = 5V, @ 10bits, [5V - 4.8mV] )
What kind of filter do we need?
  For digital domain demodulation, we need to use a band-pass filter
    No mixer, less added noise
  For analog domain demodulation, we use a simple low-pass filter
    ADC speed can be relaxed, Less Complex, Cheaper
What kind of ADC?
  Most ADCs work for 10 kHz - 1 MHz (We need 20kHz)
  With our SNR, we can achieve: ENOB = 17 bits
    For low power, SAR is suitable
  Suggest - Linear Technology LTC2386-18 (18 bit ADC)
https://www.analog.com/media/en/technical-documentation/data-sheets/238618f.pdf
We prefer RC low pass filter over LC circuits since the operating frequency is low, suitable for RC.
We suggest a Sallen-key circuit.
Output of TIA used in acquisition chain
BandPass Anti-Aliasing for Digital Domain Demodulation
Notes: Rest of the stages of the acquisition chain – A 2nd amplifier necessary? An Anti-Aliasing Filter? AD Converter? Assuming the fbias is demodulated digitally
Now what are the benefits of demodulating in analog
The TIA also acts as an LNA in a TX/RX chain thanks to Friis Eq. it should have extra low noise to reduce the noise factor up until the ADC
We prefer RC low pass filter over LC circuits since the operating frequency is low, suitable for RC.
For better performance, we suggest a sallen-Key circuit: les noise and sharper roll-off

## Slide 17: Conclusion
We achieved the best performance using the Linear Technology LTC6268
  Best TIA performance using: = 100 = 1pF, = 10kHz, = 5V
  Gain : 143 dB
  Linearity : -12.5 mdB (0.14%) @ 10KHz
  SNR Range = [148 dB, 99 dB]
  TIA NF = 1.15dB
Questions?